# Versionshinweise – DataServer Library for BioWare - 2.5.2.6397

English

Allgemeines

Systemanforderungen

Kompatibilit�t

Neue Funktionen

Bekannte Probleme

Historie

### Allgemeines

This document provides additional information to the most recent version of the software utilities for the BioWare DataServer library. This package contains the following items:

- DataServer driver installation: The installation of the MCC drivers should be performed before connecting the type 5691A or 5695A DAQ System devices. For installing the USB drivers follow the installation instructions.

- The DataServer library is intended to be integrated with external systems and requires software programming for proper integration.

- For assistance, refer to the DataServer.chm file. It contains detailed information and examples of integration.

#### Hintweise

- DataServer can be used stand-alone or together with a BioWare system.

### Systemanforderungen

#### Minimale Systemanforderungen zum Betrieb der Software

- Windows PC or Laptop

- One free USB port (if using USB Daq device)

- Windows 7 (x86 and x64), Windows 8 (x86 and x64), Windows 10 (x86 and x64)

### Kompatibilit�t

N/A

## Neue Funktionen

#### Changes and fixes for v2.5.2

- 7438: Allow 4 channel devices (5165A4) in measurement chain

### Bekannte Probleme

This section contains a list of known issues for the most recent version of the BioWare DataServer utility:

- Nil.

### Historie

#### Changes and fixes for v2.5.1

- 7438: Address minor memory leaks

#### Changes and fixes for v2.5.0

- 7429: Update calculations to provide polynomial corrected moments.

#### Changes and fixes for v2.4.12

- 7419: Disable measurement prior to setting daq mode to MODE1.

- 7421: Check/Set PTP transport layer.

- 7422: Ensure ethernet PTP sync device is configured and in-sync.

#### Changes and fixes for v2.4.11

- 7417: Add support for new Sync Device.

#### Changes and fixes for v2.4.10

- 7395: Failed to register DataServer some new Windows 11 machines.

- 7392: Update to latest MCC drivers 6.73 (Win 11 support).

#### Changes and fixes for v2.4.8

- 7364: Sampling with a rate of 2400Hz goes out of sync.

- 7363: Enable overlay of sync frame onto data for debugging.

#### Changes and fixes for v2.4.7

- 7362: External sync. fails with a small fixed-time offset

- 7353: Control panel shows multiple installed version after upgrading

- 7360: 32-bit and 64-bit versions have different build numbers

#### Changes and fixes for v2.4.6

- 7357: Using external sync rate other than 1.0 Hz fails.

- 7358: Wait 2 seconds prior to flushing trigger event queue.

- 7359: Reformat log files for easier tracking of sync events.

#### Changes and fixes for v2.4.5

- 7350: Disable Global coordinate calculations if values are set to zero.

#### Changes and fixes for v2.4.4

- 7344: Update of trigger settings on digital devices fails.

- 7345: Sync on digital devices can drift if sync pulse is missing/not captured.

- 7346: Sync. can be off by one sample on digital devices.

- 7347: Sync on falling edge throws exception.

- 7348: Disable build number increment during release builds.

#### Changes and fixes for v2.4.3

- 7338: bio-digital plate return incorrect data.

- 7337: Data acquisition reports all zeros for bio-digital plate.

- 7262: Send command to configure trigger start/stop to 'on-request'.

- 7336: Daq fails with bio-digital plate.

#### Changes and fixes for v2.4.2

- 7335: Add checks for proper ethernet device configuration

#### Changes and fixes for v2.4.1

- 7329: Update to support new Bio-Digital protocol

- 7330: Set virtual channel mode to MODE1 for Bio-Digital plates

#### Changes and fixes for v2.3.11

- 7317: Move setting rate to a parallel operation

- 7316: Resetting offsets is slow, move to parallel operation

#### Changes and fixes for v2.3.10 (interim release)

- 7313: Fix no sync / Nexus lockup (infinite looping shown in log file)

- 7312: Fix sync lockup under certain conditions (infinite looping shown in log file)

#### Changes and fixes for v2.3.9

- 7315: Add debug symbols to installer

- 7308: Change to sync on incoming pulses instead of tracking measured sync. rate

- 7309: Add version number to log file

- 7304: Add error messages if config file content is not correct for selected board type

- 7310: Add mutex to status callback (maybe fix shutdown issued)

- 7311: estimate inter sync. timing and use for improved skip/duplicate of samples

#### Changes and fixes for v2.3.8

- 7301: update plaform toolset to v142 (vs2019) and platform to 10.0

- 7302: Update check / set trigger settings

- 7303: Update error messages

- 7304: Add error messages if config file content is not correct for selected board type

- 7305: Bug fix setting PTP in labamp. Update version to 1.1

- 7306: Fix exception handling to shutdown storage thread

- 7307: update for Nike (add sync pulse info to log)

#### Changes and fixes for v2.3.7

- 7238: Improve error handling when unable to connect to Ethernet plate

#### Changes and fixes for v2.3.6

- 7203: Add support for 5th order COP correction polynomial

#### Changes and fixes for v2.3.5

- 7200: Updated XML support for motion capture location/orientation parameters.

- 7199: BugFix: Data on ethernet plates is swapped.

#### Changes and fixes for v2.2.3

- 7196: Add IDeviceMetadata interface

- 7198: Schema allows custom elements to exist in XML configuration file

#### Changes and fixes for v2.1.2

- 7194: First plate data reported for all devices

#### Changes and fixes for v2.0.1

- 7151: "Save to BioWare file" not functioning.

- 7152: Reported data stream is corrupted.

- 7148: Exception thrown when setting operate/reset on single channel devices.

- 7149: Update to work with older USB daq boxes.

- 7150: Include Ethernet samples as sub-solution.

#### Changes and fixes for v2.0.0

- 7144: Add 64-bit DataServer to system build, Include installer for 64 bit.

- 7143: Update to support Ethernet Force Plates.

- 7032: Update version of MCC on CD-ROM image to 6.71.

#### Changes and fixes for v1.6.2.2201

- 6953: Update to MCC Instacal v6.54

#### Changes and fixes for v1.6.1.2198

- 6947: Cannot get IDaqAdvanced interface from daq object in VS2015.

#### Changes and fixes for v1.6.0.2197

- 6825: Board not available message on Win 8.1 after restarting application.

- 6826: Update to MCC Instacal 6.50 drivers.

#### Changes and fixes for v1.5.1.2192

- 6583: Support USB-1616HS as a USB16161FS (no digital).

- 6443: Support MCC 6.31.

#### Changes and fixes for v1.5.0.2152

- 6443: Update to MCC Instacal drivers v6.31.

#### Changes and fixes for v1.4.0.2150

- 5883: CD-ROM contains incorrect version of MCC drivers.

- 5880: Update to MCC Instacal drivers v6.25.

#### Changes and fixes for v1.4.0.2148

- 5848: Update to MCC Instacal drivers v6.24

- 5852: Add sample files for C#.

#### Changes and fixes for v1.3.2.x

- 5547: Installation did not fully register DataServer.dll COM object for automation languages

#### Changes and fixes for v1.3.1.x

- 5434: PCIM-DAS1602 Time constant control is inverted.

- 5274: Documenation for IDevice arrVoltage() method contains a typo.

- 5289: Add method for retrieving safe array of voltages in IDaqControl2.

- 5290: Ensure channel indexing for IDevice::arrVoltages is based on the device channels.

- 5268: 5695A with external pacer clock - overrun halts system but does not issue error message.

- 5292: Add error messages when accessing data no longer in the buffer.

#### Changes and fixes for v1.3.0.x

- 3981: 5691 DAQ Box (USB-1616FS) hangs reading offsets. Requires firmware update to USB-1616FS.

- 5004: Update to support MCC Instacal Version 6.22.

- 3428: C++ Test programs have a compiler error.

#### Changes and fixes for v1.2.0.x

- 4713: Update to support MCC Instacal Version 6.20

- 4798: Startup problem when DAQ-5695 is not connected / powered-off.

- 4797: Support additional DAQ gain settings for 5695.

#### Changes and fixes for v1.0.4

- 4252 Add Pretriggering

- 4250 Modify log file options

#### Changes and fixes for v1.0.3

- DAQ 5695 - Acquiring non-sequential plate fails

#### Changes and fixes for v1.0.2

- Minor interface changes to support DAQ 5695 in BioWare

#### Changes and fixes for v1.0.1

- m_DAQ->Running reports 0 when measuring is in progress.

- Offsets are always read at +/-5V range

#### Changes and fixes for v1.0.0

- Support added for Kistler64 channel 5695 DAQ System.

- Fixed issue with Offsets ot working properly.

- Removed reading offsets hen Measure is enabled.  Requires client to manually call ReadOffsets. This was done to allow changing between measure on and measure off while data acquisition is running.

- ReadOffsets method changed to take A/D Range as a parameter.  Usually has the same value as the range used in the Start method

#### Changes and fixes for v0.9.8

- Initial version.

#### Copyright © 2021, Kistler Group. All rights reserved.